## XIGA 1600-2500nm
| NIST calibration | Linearity worst case 1000nm gain 6 | Sphere Cal repeatability | Sphere cal reproducability | measurement noise | sphere uniformity @850nm | XIGA-002 total uncertainty | end of sheet
Wavelength [nm] | % Uncertainty (k=2) | % uncertainty (k=2) | % uncertainty (2-sigma) | % uncertainty | 4-second stability (2-sigma average) | 2-sigma | RSS
1594.7039 | 0.68237 | 0.1 | 0.40586 | 0.25051 | 0.02 | 0.26 | 0.87813
1601.0464 | 0.68209 | 0.1 | 0.076832 | 0.29672 | 0.02 | 0.26 | 0.79824
1605.8605 | 0.67666 | 0.1 | 0.11447 | 0.58006 | 0.02 | 0.26 | 0.94098
1610.6545 | 0.6766 | 0.1 | 0.10879 | 0.41151 | 0.02 | 0.26 | 0.84673
1613.9004 | 0.67662 | 0.1 | 0.029104 | 0.43159 | 0.02 | 0.26 | 0.85026
1620.4401 | 0.67676 | 0.1 | 0.062607 | 0.3571 | 0.02 | 0.26 | 0.81697
1627.017 | 0.67658 | 0.1 | 0.12039 | 0.2564 | 0.02 | 0.26 | 0.78485
1630.3433 | 0.67656 | 0.1 | 0.010361 | 0.30366 | 0.02 | 0.26 | 0.7925
1635.3459 | 0.67663 | 0.1 | 0.006866 | 0.3001 | 0.02 | 0.26 | 0.79116
1638.658 | 0.67651 | 0.1 | 0.096231 | 0.32146 | 0.02 | 0.26 | 0.80515
1643.6839 | 0.67686 | 0.1 | 0.17411 | 0.41801 | 0.02 | 0.26 | 0.86092
1648.7679 | 0.67653 | 0.1 | 0.060637 | 0.25657 | 0.02 | 0.26 | 0.77794
1653.8947 | 0.67678 | 0.1 | 0.12174 | 0.24526 | 0.02 | 0.26 | 0.78167
1660.7283 | 0.6766 | 0.1 | 0.12489 | 0.33179 | 0.02 | 0.26 | 0.8133
1665.8964 | 0.67683 | 0.1 | 0.15688 | 0.16945 | 0.02 | 0.26 | 0.76774
1671.0881 | 0.6765 | 0.1 | 0.10927 | 0.16547 | 0.02 | 0.26 | 0.75827
1674.6078 | 0.67655 | 0.1 | 0.16742 | 0.21611 | 0.02 | 0.26 | 0.78132
1679.8736 | 0.67665 | 0.1 | 0.046681 | 0.13814 | 0.02 | 0.26 | 0.7464
1683.3985 | 0.67696 | 0.1 | 0.049078 | 0.1867 | 0.02 | 0.26 | 0.75732
1688.6866 | 0.67651 | 0.1 | 0.15209 | 0.25964 | 0.02 | 0.26 | 0.79134
1695.8482 | 0.67677 | 0.1 | 0.027201 | 0.34399 | 0.02 | 0.26 | 0.80938
1699.4473 | 0.6766 | 0.1 | 0.057706 | 0.55592 | 0.02 | 0.26 | 0.92096
1703.0005 | 0.67651 | 0.1 | 0.033789 | 0.15297 | 0.02 | 0.26 | 0.74847
1710.2647 | 0.67673 | 0.1 | 0.038548 | 0.29468 | 0.02 | 0.26 | 0.79012
1713.9511 | 0.67739 | 0.1 | 0.14861 | 0.25895 | 0.02 | 0.26 | 0.7912
1721.2673 | 0.67667 | 0.1 | 0.082964 | 0.22789 | 0.02 | 0.26 | 0.77117
1724.9675 | 0.67656 | 0.1 | 0.13133 | 0.35414 | 0.02 | 0.26 | 0.82364
1730.5481 | 0.67658 | 0.1 | 0.19103 | 0.41918 | 0.02 | 0.26 | 0.86485
1734.2766 | 0.67673 | 0.1 | 0.25241 | 0.38225 | 0.02 | 0.26 | 0.86359
1738.009 | 0.67661 | 0.1 | 0.099943 | 0.4033 | 0.02 | 0.26 | 0.84169
1745.5689 | 0.67666 | 0.1 | 0.12558 | 0.41721 | 0.02 | 0.26 | 0.85188
1749.3907 | 0.67663 | 0.1 | 0.071281 | 0.40667 | 0.02 | 0.26 | 0.84041
1753.2198 | 0.67655 | 0.1 | 0.084587 | 0.41131 | 0.02 | 0.26 | 0.84383
1760.9388 | 0.67659 | 0.1 | 0.17434 | 0.37573 | 0.02 | 0.26 | 0.84104
1764.8144 | 0.67707 | 0.1 | 0.18033 | 0.3072 | 0.02 | 0.26 | 0.81444
1768.6974 | 0.67681 | 0.1 | 0.027818 | 0.18653 | 0.02 | 0.26 | 0.75607
1770.6587 | 0.67686 | 0.1 | 0.12039 | 0.10302 | 0.02 | 0.26 | 0.74917
1776.5493 | 0.67665 | 0.1 | 0.19822 | 0.40539 | 0.02 | 0.26 | 0.85993
1780.4819 | 0.67666 | 0.1 | 0.12627 | 0.17008 | 0.02 | 0.26 | 0.76207
1782.4538 | 0.67681 | 0.1 | 0.09292 | 0.050131 | 0.02 | 0.26 | 0.73974
1784.4089 | 0.67688 | 0.1 | 0.037395 | 0.097955 | 0.02 | 0.26 | 0.7397
1790.383 | 0.67701 | 0.1 | 0.17979 | 0.72883 | 0.02 | 0.26 | 1.0487
1794.382 | 0.67663 | 0.1 | 0.2109 | 1.0284 | 0.02 | 0.26 | 1.2798
1798.4206 | 0.67659 | 0.1 | 0.30776 | 0.27374 | 0.02 | 0.26 | 0.83989
1804.451 | 0.67671 | 0.1 | 0.16245 | 0.19048 | 0.02 | 0.26 | 0.7737
1808.5689 | 0.67683 | 0.1 | 0.097314 | 0.29234 | 0.02 | 0.26 | 0.79438
1814.6628 | 0.6774 | 0.1 | 0.085531 | 0.4305 | 0.02 | 0.26 | 0.85412
1818.8126 | 0.67715 | 0.1 | 0.51726 | 0.33787 | 0.02 | 0.26 | 0.95825
1824.9884 | 0.67692 | 0.1 | 0.38802 | 0.3829 | 0.02 | 0.26 | 0.9129
1829.1587 | 0.67712 | 0.1 | 0.51463 | 0.29705 | 0.02 | 0.26 | 0.94317
1833.306 | 0.67661 | 0.1 | 0.83085 | 0.10853 | 0.02 | 0.26 | 1.1126
1837.5214 | 0.67748 | 0.1 | 1.1373 | 1.3906 | 0.02 | 0.26 | 1.9402
1843.8606 | 0.69562 | 0.1 | 1.1012 | 1.8475 | 0.02 | 0.26 | 2.2777
1852.3549 | 0.69594 | 0.1 | 1.0998 | 1.3941 | 0.02 | 0.26 | 1.9276
1854.5024 | 0.70676 | 0.1 | 0.47451 | 0.87401 | 0.02 | 0.26 | 1.2516
1860.9515 | 0.69571 | 0.1 | 0.75929 | 1.1887 | 0.02 | 0.26 | 1.5974
1865.2577 | 0.69833 | 0.1 | 0.79018 | 0.76394 | 0.02 | 0.26 | 1.3318
1869.6093 | 0.69578 | 0.1 | 0.5664 | 1.4247 | 0.02 | 0.26 | 1.7066
1873.9824 | 0.69566 | 0.1 | 1.1234 | 2.3682 | 0.02 | 0.26 | 2.7262
1878.364 | 0.69559 | 0.1 | 0.67519 | 0.1347 | 0.02 | 0.26 | 1.0178
1884.9642 | 0.69591 | 0.1 | 0.24671 | 0.76503 | 0.02 | 0.26 | 1.0993
1891.6318 | 0.73504 | 0.1 | 0.59564 | 0.06626 | 0.02 | 0.26 | 0.98867
1896.1189 | 0.73759 | 0.1 | 0.21135 | 0.35015 | 0.02 | 0.26 | 0.88843
1898.3641 | 0.73705 | 0.1 | 0.51016 | 0.7502 | 0.02 | 0.26 | 1.2018
1902.8754 | 0.83307 | 0.1 | 0.67012 | 0.96979 | 0.02 | 0.26 | 1.4702
1907.3673 | 0.77691 | 0.1 | 0.94558 | 0.85937 | 0.02 | 0.26 | 1.5213
1909.6287 | 0.71895 | 0.1 | 0.98575 | 0.4164 | 0.02 | 0.26 | 1.3191
1916.5072 | 0.71888 | 0.1 | 0.73862 | 0.3111 | 0.02 | 0.26 | 1.1123
1921.0608 | 0.71976 | 0.1 | 0.55638 | 0.95541 | 0.02 | 0.26 | 1.3485
1925.6935 | 0.71851 | 0.1 | 0.19715 | 0.61552 | 0.02 | 0.26 | 1.006
1932.6244 | 0.72036 | 0.1 | 0.26467 | 0.35082 | 0.02 | 0.26 | 0.88884
1937.2835 | 0.72042 | 0.1 | 0.22686 | 0.23439 | 0.02 | 0.26 | 0.8387
1941.9999 | 0.71957 | 0.1 | 0.060527 | 0.16454 | 0.02 | 0.26 | 0.79153
1946.687 | 0.72878 | 0.1 | 0.19047 | 0.47175 | 0.02 | 0.26 | 0.93164
1951.4016 | 0.71966 | 0.1 | 0.16025 | 0.33508 | 0.02 | 0.26 | 0.85666
1956.2042 | 0.72103 | 0.1 | 0.16797 | 0.26146 | 0.02 | 0.26 | 0.83335
1963.3314 | 0.72272 | 0.1 | 0.25634 | 0.30693 | 0.02 | 0.26 | 0.87192
1973.0225 | 0.71965 | 0.1 | 0.49064 | 0.38634 | 0.02 | 0.26 | 0.99291
1980.2568 | 0.72025 | 0.1 | 0.93736 | 0.83719 | 0.02 | 0.26 | 1.4752
1987.57 | 0.90392 | 0.1 | 0.085957 | 1.0707 | 0.02 | 0.26 | 1.4314
1994.84 | 0.7559 | 0.1 | 0.8464 | 1.0501 | 0.02 | 0.26 | 1.5711
2001.77 | 0.85131 | 0.1 | 0.75375 | 0.27604 | 0.02 | 0.26 | 1.2029
2009.5 | 0.79565 | 0.1 | 0.47349 | 0.35892 | 0.02 | 0.26 | 1.0315
2016.56 | 0.81069 | 0.1 | 0.046572 | 0.4511 | 0.02 | 0.26 | 0.96999
2022.87 | 0.77234 | 0.1 | 0.42559 | 0.31147 | 0.02 | 0.26 | 0.97604
2026.82 | 0.92439 | 0.1 | 0.34579 | 0.19813 | 0.02 | 0.26 | 1.0447
2029.97 | 0.8474 | 0.1 | 0.66885 | 0.24472 | 0.02 | 0.26 | 1.1416
2036.82 | 0.81523 | 0.1 | 0.36542 | 0.12647 | 0.02 | 0.26 | 0.94453
2044.04 | 0.85515 | 0.1 | 0.05958 | 0.76365 | 0.02 | 0.26 | 1.1815
2050.94 | 0.76994 | 0.1 | 0.94911 | 0.63353 | 0.02 | 0.26 | 1.4046
2058.64 | 0.77641 | 0.1 | 0.56791 | 0.68627 | 0.02 | 0.26 | 1.2142
2065.65 | 0.80266 | 0.1 | 0.38662 | 0.021474 | 0.02 | 0.26 | 0.93391
2072.3 | 0.83855 | 0.1 | 0.26504 | 0.33523 | 0.02 | 0.26 | 0.98173
2079.27 | 0.78194 | 0.1 | 0.37708 | 0.23298 | 0.02 | 0.26 | 0.94122
2086.15 | 0.8182 | 0.1 | 0.52504 | 0.37029 | 0.02 | 0.26 | 1.0771
2092.9 | 0.81652 | 0.1 | 0.63684 | 0.39535 | 0.02 | 0.26 | 1.1431
2099.77 | 0.79668 | 0.1 | 0.20604 | 0.28836 | 0.02 | 0.26 | 0.91558
2106.55 | 0.8615 | 0.1 | 0.51638 | 0.54235 | 0.02 | 0.26 | 1.1752
2113.45 | 0.78694 | 0.1 | 0.53979 | 0.10639 | 0.02 | 0.26 | 0.99998
2120.33 | 0.75414 | 0.1 | 0.38189 | 0.3303 | 0.02 | 0.26 | 0.94957
2127 | 0.98492 | 0.1 | 0.47 | 0.35065 | 0.02 | 0.26 | 1.1798
2134.02 | 0.81956 | 0.1 | 0.39192 | 0.36478 | 0.02 | 0.26 | 1.018
2140.56 | 0.79826 | 0.1 | 0.507 | 0.18432 | 0.02 | 0.26 | 1.0031
2147.39 | 0.78976 | 0.1 | 0.63047 | 0.28207 | 0.02 | 0.26 | 1.0857
2154.15 | 0.77668 | 0.1 | 1.0664 | 0.01346 | 0.02 | 0.26 | 1.3485
2160.71 | 0.81954 | 0.1 | 0.31171 | 0.58915 | 0.02 | 0.26 | 1.0927
2167.36 | 0.80438 | 0.1 | 0.56506 | 0.50937 | 0.02 | 0.26 | 1.1418
2174.22 | 0.77546 | 0.1 | 0.11745 | 0.32319 | 0.02 | 0.26 | 0.89308
2180.74 | 0.79345 | 0.1 | 0.41766 | 0.38259 | 0.02 | 0.26 | 1.0141
2187.3 | 0.8249 | 0.1 | 0.73832 | 0.44193 | 0.02 | 0.26 | 1.2243
2193.84 | 0.80529 | 0.1 | 1.0338 | 0.34945 | 0.02 | 0.26 | 1.3847
2200.69 | 0.81464 | 0.1 | 0.25143 | 0.26398 | 0.02 | 0.26 | 0.93517
2207.11 | 0.79259 | 0.1 | 0.18384 | 0.36146 | 0.02 | 0.26 | 0.93308
2213.83 | 0.72584 | 0.1 | 0.13491 | 0.37154 | 0.02 | 0.26 | 0.8724
2220.53 | 0.76047 | 0.1 | 0.26365 | 0.52011 | 0.02 | 0.26 | 0.99817
2226.84 | 0.74645 | 0.1 | 0.35607 | 0.17396 | 0.02 | 0.26 | 0.89008
2233.66 | 0.73888 | 0.1 | 0.31758 | 0.14476 | 0.02 | 0.26 | 0.86357
2240.05 | 0.78219 | 0.1 | 0.10791 | 0.36119 | 0.02 | 0.26 | 0.9121
2246.5 | 0.73935 | 0.1 | 0.16443 | 0.19823 | 0.02 | 0.26 | 0.83124
2252.97 | 0.81097 | 0.1 | 0.1479 | 0.44122 | 0.02 | 0.26 | 0.97582
2259.46 | 0.74479 | 0.1 | 0.2379 | 0.31588 | 0.02 | 0.26 | 0.88831
2266 | 0.77542 | 0.1 | 0.18636 | 0.44669 | 0.02 | 0.26 | 0.95579
2272.38 | 0.75869 | 0.1 | 0.041787 | 0.31471 | 0.02 | 0.26 | 0.86856
2278.74 | 0.88083 | 0.1 | 0.18578 | 0.40749 | 0.02 | 0.26 | 1.0268
2285.39 | 0.80954 | 0.1 | 0.23017 | 0.12254 | 0.02 | 0.26 | 0.89519
2291.61 | 0.78633 | 0.1 | 0.12792 | 0.23331 | 0.02 | 0.26 | 0.87585
2298.01 | 0.73143 | 0.1 | 0.088514 | 0.18351 | 0.02 | 0.26 | 0.80902
2304.29 | 0.78262 | 0.1 | 0.23598 | 0.27168 | 0.02 | 0.26 | 0.90553
2310.95 | 0.7645 | 0.1 | 0.22832 | 0.63512 | 0.02 | 0.26 | 1.0573
2317.18 | 0.79554 | 0.1 | 0.44175 | 0.38442 | 0.02 | 0.26 | 1.0266
2323.25 | 0.8039 | 0.1 | 0.071017 | 1.0559 | 0.02 | 0.26 | 1.358
2329.59 | 0.76156 | 0.1 | 0.22858 | 0.29041 | 0.02 | 0.26 | 0.89138
2336.07 | 0.78263 | 0.1 | 0.13588 | 0.41796 | 0.02 | 0.26 | 0.94004
2342.49 | 0.77668 | 0.1 | 0.21977 | 0.22559 | 0.02 | 0.26 | 0.88342
2348.58 | 0.91138 | 0.1 | 0.29048 | 0.50706 | 0.02 | 0.26 | 1.1181
2354.81 | 0.7613 | 0.1 | 0.2599 | 0.10243 | 0.02 | 0.26 | 0.85768
2361.19 | 0.79513 | 0.1 | 0.27653 | 0.046062 | 0.02 | 0.26 | 0.88816
2367.31 | 0.79377 | 0.1 | 0.4243 | 0.12603 | 0.02 | 0.26 | 0.95079
2373.47 | 0.78805 | 0.1 | 0.43636 | 0.4168 | 0.02 | 0.26 | 1.0311
2379.64 | 0.85391 | 0.1 | 0.36078 | 0.92975 | 0.02 | 0.26 | 1.3423
2385.85 | 0.87233 | 0.1 | 0.25627 | 0.25803 | 0.02 | 0.26 | 0.9855
2392.12 | 0.80374 | 0.1 | 0.068915 | 0.76356 | 0.02 | 0.26 | 1.1453
2398.12 | 0.78223 | 0.1 | 0.28102 | 0.36466 | 0.02 | 0.26 | 0.94965
2404.19 | 0.88622 | 0.1 | 0.74609 | 0.96118 | 0.02 | 0.26 | 1.531
2410.54 | 0.75714 | 0.1 | 0.53969 | 0.19341 | 0.02 | 0.26 | 0.98992
2416.35 | 0.76353 | 0.1 | 0.35116 | 0.508 | 0.02 | 0.26 | 1.021
2422.78 | 0.76696 | 0.1 | 0.33629 | 0.57183 | 0.02 | 0.26 | 1.0518
2429.23 | 0.92355 | 0.1 | 0.16785 | 0.667 | 0.02 | 0.26 | 1.1849
2435.17 | 0.88416 | 0.1 | 0.86597 | 0.78611 | 0.02 | 0.26 | 1.4925
2441.11 | 0.74932 | 0.1 | 0.3934 | 0.53788 | 0.02 | 0.26 | 1.0409
2447.09 | 0.97488 | 0.1 | 0.1891 | 0.69111 | 0.02 | 0.26 | 1.2417
2453.12 | 0.86607 | 0.1 | 0.4076 | 0.41349 | 0.02 | 0.26 | 1.0794
2459.18 | 0.79938 | 0.1 | 0.44064 | 0.51876 | 0.02 | 0.26 | 1.0864
2465.26 | 0.82237 | 0.1 | 0.46023 | 0.43747 | 0.02 | 0.26 | 1.0759
2471.3 | 0.86744 | 0.1 | 0.28507 | 1.0898 | 0.02 | 0.26 | 1.449
2476.36 | 0.91609 | 0.1 | 0.15724 | 0.6781 | 0.02 | 0.26 | 1.184
2482.54 | 0.78494 | 0.1 | 0.36901 | 0.54772 | 0.02 | 0.26 | 1.0631
2488.17 | 0.75508 | 0.1 | 0.17056 | 0.52155 | 0.02 | 0.26 | 0.97429
2494.39 | 0.81452 | 0.1 | 0.41175 | 0.99239 | 0.02 | 0.26 | 1.3769
2500.09 | 0.78132 | 0.1 | 0.8006 | 1.9853 | 0.02 | 0.26 | 2.2958
2505.81 | 0.86741 | 0.1 | 0.8006 | 1.9853 | 0.02 | 0.26 | 2.3265

## IGA 950-1600nm
| NIST calibration | Linearity worst case 1000nm gain 6 | Sphere Cal repeatability | Sphere cal reproducability | measurement noise | sphere uniformity @ 850nm | DET-8-IGA-109 total uncertainty | end of data
Wavelength [nm] | % uncertainty (k=2) | % uncertainty (k=2) | % uncertainty (2-sigma) | % uncertainty | 4-second stability (2-sigma average) | 2-sigma | RSS
945.5312 | 0.71975 | 0.022 | 0.029128 | 0.28794 | 0.02 | 0.26 | 0.81871
950.5741 | 0.71982 | 0.022 | 0.029128 | 0.28794 | 0.02 | 0.26 | 0.81877
955.1184 | 0.71973 | 0.022 | 0.029128 | 0.28794 | 0.02 | 0.26 | 0.81869
959.6974 | 0.71974 | 0.022 | 0.029128 | 0.27635 | 0.02 | 0.26 | 0.8147
964.904 | 0.71973 | 0.022 | 0.029128 | 0.26353 | 0.02 | 0.26 | 0.81043
969.5757 | 0.71974 | 0.022 | 0.029128 | 0.23085 | 0.02 | 0.26 | 0.8004
975.4943 | 0.71975 | 0.022 | 0.029128 | 0.22521 | 0.02 | 0.26 | 0.7988
979.6617 | 0.71975 | 0.022 | 0.029128 | 0.20689 | 0.02 | 0.26 | 0.79384
984.5026 | 0.71973 | 0.022 | 0.029128 | 0.21655 | 0.02 | 0.26 | 0.79639
990.5905 | 0.71973 | 0.022 | 0.029128 | 0.20976 | 0.02 | 0.26 | 0.79457
994.8974 | 0.71975 | 0.022 | 0.029128 | 0.20409 | 0.02 | 0.26 | 0.79311
999.8773 | 0.7198 | 0.022 | 0.029128 | 0.18302 | 0.02 | 0.26 | 0.788
1005.5282 | 0.6821 | 0.022 | 0.014823 | 0.18242 | 0.02 | 0.26 | 0.75316
1010.5994 | 0.68214 | 0.022 | 0.023031 | 0.16618 | 0.02 | 0.26 | 0.74963
1014.4487 | 0.68209 | 0.022 | 0.022867 | 0.15596 | 0.02 | 0.26 | 0.74738
1020.2664 | 0.68216 | 0.022 | 0.017927 | 0.15397 | 0.02 | 0.26 | 0.7469
1025.5138 | 0.68209 | 0.022 | 0.019979 | 0.11004 | 0.02 | 0.26 | 0.73908
1029.4881 | 0.68213 | 0.022 | 0.063297 | 0.10566 | 0.02 | 0.26 | 0.74092
1032.1341 | 0.68219 | 0.022 | 0.053265 | 0.13209 | 0.02 | 0.26 | 0.74441
1034.8072 | 0.68207 | 0.022 | 0.01373 | 0.1461 | 0.02 | 0.26 | 0.74514
1040.1926 | 0.68209 | 0.022 | 0.0060024 | 0.17307 | 0.02 | 0.26 | 0.75081
1045.6198 | 0.68211 | 0.022 | 0.05636 | 0.16826 | 0.02 | 0.26 | 0.75183
1048.3724 | 0.70137 | 0.022 | 0.13841 | 0.16224 | 0.02 | 0.26 | 0.77838
1049.4886 | 0.70306 | 0.022 | 0.1465 | 0.15857 | 0.02 | 0.26 | 0.78063
1051.1253 | 0.70615 | 0.022 | 0.12381 | 0.15096 | 0.02 | 0.26 | 0.77798
1054.5853 | 0.70504 | 0.022 | 0.22193 | 0.15358 | 0.02 | 0.26 | 0.799
1055.1258 | 0.70114 | 0.022 | 0.28764 | 0.14489 | 0.02 | 0.26 | 0.81475
1058.0763 | 0.70421 | 0.022 | 0.35825 | 0.13912 | 0.02 | 0.26 | 0.84386
1060.6172 | 0.70327 | 0.022 | 0.30838 | 0.12642 | 0.02 | 0.26 | 0.82107
1062.3112 | 0.70781 | 0.022 | 0.21817 | 0.13339 | 0.02 | 0.26 | 0.79679
1074.4597 | 0.70273 | 0.022 | 0.035036 | 0.14351 | 0.02 | 0.26 | 0.76429
1079.5233 | 0.71143 | 0.022 | 0.019352 | 0.17382 | 0.02 | 0.26 | 0.77795
1085.3977 | 0.71086 | 0.022 | 0.016152 | 0.18648 | 0.02 | 0.26 | 0.78028
1091.3263 | 0.70127 | 0.022 | 0.059151 | 0.13814 | 0.02 | 0.26 | 0.76345
1094.3184 | 0.7052 | 0.022 | 0.042827 | 0.12729 | 0.02 | 0.26 | 0.76408
1099.6302 | 0.7068 | 0.022 | 0.082962 | 0.14565 | 0.02 | 0.26 | 0.7721
1103.5397 | 0.70434 | 0.022 | 0.064305 | 0.12566 | 0.02 | 0.26 | 0.76452
1110.3894 | 0.70206 | 0.022 | 0.037616 | 0.097382 | 0.02 | 0.26 | 0.75649
1114.2097 | 0.68401 | 0.022 | 0.036784 | 0.096311 | 0.02 | 0.26 | 0.73958
1120.4088 | 0.68586 | 0.022 | 0.092138 | 0.11667 | 0.02 | 0.26 | 0.74899
1124.3418 | 0.69412 | 0.022 | 0.07929 | 0.12048 | 0.02 | 0.26 | 0.7557
1130.7374 | 0.69036 | 0.022 | 0.06459 | 0.088442 | 0.02 | 0.26 | 0.74637
1136.4509 | 0.69066 | 0.022 | 0.12225 | 0.19441 | 0.02 | 0.26 | 0.77345
1138.8869 | 0.68465 | 0.022 | 0.086178 | 0.16989 | 0.02 | 0.26 | 0.75731
1145.3531 | 0.69342 | 0.022 | 0.033135 | 0.089406 | 0.02 | 0.26 | 0.74726
1150.2481 | 0.68226 | 0.022 | 0.079499 | 0.14289 | 0.02 | 0.26 | 0.7488
1154.3894 | 0.71618 | 0.022 | 0.024135 | 0.17543 | 0.02 | 0.26 | 0.78279
1158.6023 | 0.68268 | 0.022 | 0.018142 | 0.079209 | 0.02 | 0.26 | 0.73562
1163.6523 | 0.84431 | 0.022 | 0.02668 | 0.16211 | 0.02 | 0.26 | 0.89908
1168.7521 | 0.83431 | 0.022 | 0.022194 | 0.14067 | 0.02 | 0.26 | 0.88591
1173.8903 | 0.83408 | 0.022 | 0.047798 | 0.1138 | 0.02 | 0.26 | 0.88284
1180.8241 | 0.83416 | 0.022 | 0.033186 | 0.15356 | 0.02 | 0.26 | 0.88825
1186.0705 | 0.70089 | 0.022 | 0.04875 | 0.16848 | 0.02 | 0.26 | 0.76843
1188.7101 | 0.70089 | 0.022 | 0.05853 | 0.17117 | 0.02 | 0.26 | 0.76971
1194.0242 | 0.68208 | 0.022 | 0.057145 | 0.18776 | 0.02 | 0.26 | 0.75646
1199.3927 | 0.68218 | 0.022 | 0.046162 | 0.15681 | 0.02 | 0.26 | 0.74871
1203.9072 | 0.68209 | 0.022 | 0.051109 | 0.18859 | 0.02 | 0.26 | 0.75625
1211.1805 | 0.68209 | 0.022 | 0.067179 | 0.14314 | 0.02 | 0.26 | 0.74748
1214.8601 | 0.68208 | 0.022 | 0.054462 | 0.15715 | 0.02 | 0.26 | 0.74925
1220.4234 | 0.68207 | 0.022 | 0.046429 | 0.20053 | 0.02 | 0.26 | 0.75899
1225.1021 | 0.68208 | 0.022 | 0.033081 | 0.20752 | 0.02 | 0.26 | 0.76018
1228.866 | 0.68208 | 0.022 | 0.042882 | 0.23251 | 0.02 | 0.26 | 0.76787
1234.5683 | 0.68215 | 0.022 | 0.062715 | 0.18829 | 0.02 | 0.26 | 0.7571
1238.3788 | 0.6821 | 0.022 | 0.059515 | 0.19073 | 0.02 | 0.26 | 0.75741
1244.1528 | 0.68209 | 0.022 | 0.015385 | 0.18851 | 0.02 | 0.26 | 0.75465
1250.0095 | 0.70092 | 0.022 | 0.013422 | 0.17772 | 0.02 | 0.26 | 0.76911
1255.874 | 0.68213 | 0.022 | 0.10824 | 0.20637 | 0.02 | 0.26 | 0.76687
1259.8266 | 0.68211 | 0.022 | 0.10422 | 0.13919 | 0.02 | 0.26 | 0.751
1265.8216 | 0.68211 | 0.022 | 0.11684 | 0.13226 | 0.02 | 0.26 | 0.7516
1269.8192 | 0.6821 | 0.022 | 0.12164 | 0.12591 | 0.02 | 0.26 | 0.75126
1274.8984 | 0.68209 | 0.022 | 0.11692 | 0.15502 | 0.02 | 0.26 | 0.75593
1278.9568 | 0.68212 | 0.022 | 0.11587 | 0.15379 | 0.02 | 0.26 | 0.75554
1285.1335 | 0.68211 | 0.022 | 0.10684 | 0.16879 | 0.02 | 0.26 | 0.7574
1290.3078 | 0.68207 | 0.022 | 0.11247 | 0.18267 | 0.02 | 0.26 | 0.7614
1295.5198 | 0.68209 | 0.022 | 0.093362 | 0.16521 | 0.02 | 0.26 | 0.75481
1300.7883 | 0.68209 | 0.022 | 0.082174 | 0.1465 | 0.02 | 0.26 | 0.74963
1305.0285 | 0.68209 | 0.022 | 0.084548 | 0.17742 | 0.02 | 0.26 | 0.75654
1310.3678 | 0.6821 | 0.022 | 0.082157 | 0.16522 | 0.02 | 0.26 | 0.75352
1314.6599 | 0.68212 | 0.022 | 0.071144 | 0.15523 | 0.02 | 0.26 | 0.75029
1320.0996 | 0.6821 | 0.022 | 0.081787 | 0.089315 | 0.02 | 0.26 | 0.74055
1324.4794 | 0.68208 | 0.022 | 0.070446 | 0.084858 | 0.02 | 0.26 | 0.73884
1328.8635 | 0.68208 | 0.022 | 0.073347 | 0.12309 | 0.02 | 0.26 | 0.74448
1335.4974 | 0.6821 | 0.022 | 0.052716 | 0.0946 | 0.02 | 0.26 | 0.73857
1339.9594 | 0.6821 | 0.022 | 0.058431 | 0.12628 | 0.02 | 0.26 | 0.74371
1344.4626 | 0.68208 | 0.022 | 0.057889 | 0.16984 | 0.02 | 0.26 | 0.75227
1348.991 | 0.68208 | 0.022 | 0.03992 | 0.17125 | 0.02 | 0.26 | 0.75143
1355.8464 | 0.68213 | 0.022 | 0.25267 | 0.469 | 0.02 | 0.26 | 0.9042
1359.2926 | 0.68209 | 0.022 | 0.061595 | 0.44702 | 0.02 | 0.26 | 0.85869
1365.1008 | 0.68304 | 0.022 | 0.046262 | 0.84384 | 0.02 | 0.26 | 1.1177
1369.7662 | 0.6821 | 0.022 | 0.17949 | 0.43131 | 0.02 | 0.26 | 0.86718
1374.4524 | 0.68226 | 0.022 | 0.14667 | 0.34243 | 0.02 | 0.26 | 0.82021
1379.1724 | 0.68212 | 0.022 | 0.2512 | 0.59446 | 0.02 | 0.26 | 0.97481
1383.935 | 0.68215 | 0.022 | 0.39934 | 0.92684 | 0.02 | 0.26 | 1.2459
1388.7236 | 0.68212 | 0.022 | 0.29126 | 0.24802 | 0.02 | 0.26 | 0.82469
1393.5754 | 0.68208 | 0.022 | 0.23248 | 0.42966 | 0.02 | 0.26 | 0.87885
1400.8997 | 0.68213 | 0.022 | 0.40836 | 0.54384 | 0.02 | 0.26 | 0.99815
1405.8182 | 0.68213 | 0.022 | 0.21674 | 0.42876 | 0.02 | 0.26 | 0.87442
1410.7748 | 0.68211 | 0.022 | 0.081654 | 0.20911 | 0.02 | 0.26 | 0.76429
1414.5383 | 0.68216 | 0.022 | 0.11057 | 0.63086 | 0.02 | 0.26 | 0.97161
1419.5455 | 0.68211 | 0.022 | 0.067257 | 0.19127 | 0.02 | 0.26 | 0.7582
1424.5932 | 0.68219 | 0.022 | 0.069652 | 0.19011 | 0.02 | 0.26 | 0.75819
1429.6725 | 0.68211 | 0.022 | 0.050312 | 0.1979 | 0.02 | 0.26 | 0.75859
1434.8042 | 0.68221 | 0.022 | 0.055502 | 0.2471 | 0.02 | 0.26 | 0.77332
1438.6578 | 0.68224 | 0.022 | 0.05376 | 0.091914 | 0.02 | 0.26 | 0.73843
1443.8532 | 0.68219 | 0.022 | 0.080555 | 0.20371 | 0.02 | 0.26 | 0.7628
1449.0762 | 0.68236 | 0.022 | 0.18646 | 0.39754 | 0.02 | 0.26 | 0.85259
1454.3276 | 0.68228 | 0.022 | 0.10203 | 0.24781 | 0.02 | 0.26 | 0.77834
1459.6294 | 0.6822 | 0.022 | 0.027046 | 0.19592 | 0.02 | 0.26 | 0.75696
1464.9724 | 0.68217 | 0.022 | 0.028183 | 0.12221 | 0.02 | 0.26 | 0.74133
1470.3341 | 0.68222 | 0.022 | 0.059512 | 0.30415 | 0.02 | 0.26 | 0.7937
1475.7449 | 0.68214 | 0.022 | 0.056819 | 0.11686 | 0.02 | 0.26 | 0.74208
1478.481 | 0.68219 | 0.022 | 0.074744 | 0.091408 | 0.02 | 0.26 | 0.74014
1485.3301 | 0.68227 | 0.022 | 0.038991 | 0.16979 | 0.02 | 0.26 | 0.75122
1489.4676 | 0.68215 | 0.022 | 0.032219 | 0.19861 | 0.02 | 0.26 | 0.75783
1495.0308 | 0.68216 | 0.022 | 0.042839 | 0.18844 | 0.02 | 0.26 | 0.75576
1499.2264 | 0.68208 | 0.022 | 0.029482 | 0.18567 | 0.02 | 0.26 | 0.75436
1504.8628 | 0.68215 | 0.022 | 0.011703 | 0.21036 | 0.02 | 0.26 | 0.76039
1510.5401 | 0.68211 | 0.022 | 0.029042 | 0.22065 | 0.02 | 0.26 | 0.76373
1513.3984 | 0.68211 | 0.022 | 0.021022 | 0.22526 | 0.02 | 0.26 | 0.76482
1519.1451 | 0.68213 | 0.022 | 0.018522 | 0.20671 | 0.02 | 0.26 | 0.75951
1524.9147 | 0.68217 | 0.022 | 0.039358 | 0.23086 | 0.02 | 0.26 | 0.76726
1529.274 | 0.6821 | 0.022 | 0.016727 | 0.23027 | 0.02 | 0.26 | 0.76619
1535.1506 | 0.68225 | 0.022 | 0.012245 | 0.22555 | 0.02 | 0.26 | 0.76484
1538.0648 | 0.68211 | 0.022 | 0.018421 | 0.19262 | 0.02 | 0.26 | 0.75578
1544.0279 | 0.68222 | 0.022 | 0.070867 | 0.20315 | 0.02 | 0.26 | 0.76171
1549.9756 | 0.68219 | 0.022 | 0.099457 | 0.17467 | 0.02 | 0.26 | 0.7578
1554.4927 | 0.68215 | 0.022 | 0.055748 | 0.1301 | 0.02 | 0.26 | 0.74421
1560.5668 | 0.68213 | 0.022 | 0.044024 | 0.14299 | 0.02 | 0.26 | 0.74577
1563.5675 | 0.68211 | 0.022 | 0.057273 | 0.15884 | 0.02 | 0.26 | 0.74985
1569.7112 | 0.68212 | 0.022 | 0.063653 | 0.22197 | 0.02 | 0.26 | 0.76622
1575.8816 | 0.68223 | 0.022 | 0.069547 | 0.17022 | 0.02 | 0.26 | 0.75348
1578.9947 | 0.68216 | 0.022 | 0.061072 | 0.17675 | 0.02 | 0.26 | 0.75418
1585.2221 | 0.68228 | 0.022 | 0.084483 | 0.20792 | 0.02 | 0.26 | 0.76443
1588.3806 | 0.68254 | 0.022 | 0.093009 | 0.13864 | 0.02 | 0.26 | 0.74981
1594.7051 | 0.6823 | 0.022 | 0.087369 | 0.11215 | 0.02 | 0.26 | 0.74446
1601.051 | 0.67658 | 0.022 | 0.087369 | 0.070903 | 0.02 | 0.26 | 0.7341

## Si 340-950nm
| NIST calibration | Linearity (TR worst case at 850, gain=6) | Sphere Cal repeatability | Sphere Cal reproducability (GLAMR 2024-2025, based on RANGE) | measurement noise | sphere uniformity | fluorescence | LTD-11-104 total uncertainty | end of sheet
wavelength [nm] | % uncertainty (k=2) | % uncertainty (k=2) | % uncertainty (2-sigma) | % uncertainty (k=2) | 4-second stability (2-sigma average) | 2-sigma |  | RSS
340.0314 | 0.39526 | 0.041 | 0.1129 | 0.10722 | 0.02 | 0.26 | 0.0003 | 0.50015
343.7066 | 1.2942 | 0.041 | 0.1129 | 0.14303 | 0.02 | 0.26 | 0.0003 | 1.3333
348.157 | 1.3522 | 0.041 | 0.89623 | 0.49965 | 0.02 | 0.26 | 0.0003 | 1.7179
350.089 | 0.3543 | 0.041 | 0.25366 | 0.91801 | 0.02 | 0.26 | 0.0003 | 1.0499
355.509 | 0.48779 | 0.041 | 0.11857 | 0.76312 | 0.02 | 0.26 | 0.0003 | 0.95081
360.023 | 0.34339 | 0.041 | 0.63205 | 0.37498 | 0.02 | 0.26 | 0.0003 | 0.85305
365.2939 | 0.34304 | 0.041 | 0.1154 | 0.3298 | 0.02 | 0.26 | 0.0003 | 0.55627
370.0094 | 0.34019 | 0.041 | 0.43129 | 0.38487 | 0.02 | 0.26 | 0.0003 | 0.7208
375.0199 | 0.34027 | 0.041 | 0.1013 | 0.33785 | 0.02 | 0.26 | 0.0003 | 0.55666
380.227 | 0.16298 | 0.041 | 0.069737 | 0.13724 | 0.02 | 0.26 | 0.0003 | 0.34633
384.9082 | 0.16301 | 0.041 | 0.11768 | 0.26319 | 0.02 | 0.26 | 0.0003 | 0.42353
390.1437 | 0.14988 | 0.041 | 0.2287 | 0.26449 | 0.02 | 0.26 | 0.0003 | 0.46304
394.9495 | 0.17621 | 0.041 | 0.11104 | 0.36049 | 0.02 | 0.26 | 0.0003 | 0.49297
399.519 | 0.14971 | 0.041 | 0.1707 | 0.30292 | 0.02 | 0.26 | 0.0003 | 0.46151
400.0707 | 0.15005 | 0.041 | 0.54668 | 0.3133 | 0.02 | 0.26 | 0.0003 | 0.69943
404.9212 | 0.14893 | 0.041 | 0.14871 | 0.33665 | 0.02 | 0.26 | 0.0003 | 0.47677
409.8826 | 0.14339 | 0.041 | 0.0087504 | 0.30114 | 0.02 | 0.26 | 0.0003 | 0.42545
414.9801 | 0.14367 | 0.041 | 0.12583 | 0.23368 | 0.02 | 0.26 | 0.0003 | 0.40095
419.9776 | 0.1478 | 0.041 | 1.4177 | 0.19239 | 0.02 | 0.26 | 0.0003 | 1.4623
425.3297 | 0.1436 | 0.041 | 0.74983 | 0.2062 | 0.02 | 0.26 | 0.0003 | 0.83371
430.3619 | 0.14259 | 0.041 | 0.24469 | 0.22455 | 0.02 | 0.26 | 0.0003 | 0.44756
435.0315 | 0.14272 | 0.041 | 0.065783 | 0.1381 | 0.02 | 0.26 | 0.0003 | 0.33682
440.0508 | 0.14234 | 0.041 | 0.10292 | 0.12624 | 0.02 | 0.26 | 0.0003 | 0.34128
444.9384 | 0.14348 | 0.041 | 0.15769 | 0.11943 | 0.02 | 0.26 | 0.0003 | 0.35972
449.6848 | 0.20017 | 0.041 | 0.093263 | 0.13809 | 0.02 | 0.26 | 0.0003 | 0.37083
455.048 | 0.14866 | 0.041 | 0.14581 | 0.10868 | 0.02 | 0.26 | 0.0003 | 0.35335
460.2803 | 0.14204 | 0.041 | 0.1026 | 0.058228 | 0.02 | 0.26 | 0.0003 | 0.32214
464.8339 | 0.14227 | 0.041 | 0.015443 | 0.099087 | 0.02 | 0.26 | 0.0003 | 0.31619
470.2996 | 0.14282 | 0.041 | 0.061508 | 0.1592 | 0.02 | 0.26 | 0.0003 | 0.34527
475.6109 | 0.15599 | 0.041 | 0.034438 | 0.12043 | 0.02 | 0.26 | 0.0003 | 0.33121
480.1722 | 0.15684 | 0.041 | 0.11585 | 0.12633 | 0.02 | 0.26 | 0.0003 | 0.35165
485.4147 | 0.15737 | 0.041 | 0.044458 | 0.15137 | 0.02 | 0.26 | 0.0003 | 0.34545
490.1739 | 0.15688 | 0.041 | 0.062783 | 0.01865 | 0.02 | 0.26 | 0.0003 | 0.31398
495.3387 | 0.15473 | 0.041 | 0.12081 | 0.14415 | 0.02 | 0.26 | 0.0003 | 0.35916
500.2925 | 0.15445 | 0.041 | 0.2377 | 0.045564 | 0.02 | 0.26 | 0.0003 | 0.39002
504.7174 | 0.15421 | 0.041 | 0.68278 | 0.0043205 | 0.02 | 0.26 |  | 0.74811
510.1872 | 0.15402 | 0.041 | 0.14793 | 0.011058 | 0.02 | 0.26 |  | 0.33972
515.1148 | 0.15395 | 0.041 | 0.27135 | 0.13765 | 0.02 | 0.26 |  | 0.43124
519.8017 | 0.15417 | 0.041 | 0.397 | 0.051064 | 0.02 | 0.26 |  | 0.50365
524.922 | 0.15394 | 0.041 | 0.80497 | 0.039546 | 0.02 | 0.26 |  | 0.86193
530.1707 | 0.15621 | 0.041 | 1.0939 | 0.048763 | 0.02 | 0.26 |  | 1.1372
534.067 | 0.1599 | 0.041 | 0.25777 | 0.012081 | 0.02 | 0.26 |  | 0.4023
537.2717 | 0.1539 | 0.041 | 0.028483 | 0.075903 | 0.02 | 0.26 |  | 0.31613
538.892 | 0.15981 | 0.041 | 0.018662 | 0.094284 | 0.02 | 0.26 |  | 0.3232
541.2749 | 0.15375 | 0.041 | 0.21267 | 0.049823 | 0.02 | 0.26 |  | 0.37554
545.0117 | 0.15412 | 0.041 | 0.14608 | 0.047442 | 0.02 | 0.26 |  | 0.34209
549.8829 | 0.15296 | 0.041 | 0.16512 | 0.08923 | 0.02 | 0.26 |  | 0.3582
555.1863 | 0.15526 | 0.041 | 0.0064882 | 0.11649 | 0.02 | 0.26 |  | 0.32771
559.563 | 0.16013 | 0.041 | 0.024877 | 0.11267 | 0.02 | 0.26 |  | 0.3296
564.495 | 0.16164 | 0.041 | 0.01948 | 0.11808 | 0.02 | 0.26 |  | 0.33186
569.518 | 0.15948 | 0.041 | 0.037751 | 0.11621 | 0.02 | 0.26 |  | 0.33173
574.723 | 0.16056 | 0.041 | 0.025317 | 0.10557 | 0.02 | 0.26 |  | 0.32749
579.605 | 0.16021 | 0.041 | 0.010101 | 0.11683 | 0.02 | 0.26 |  | 0.3303
584.767 | 0.16049 | 0.041 | 0.017202 | 0.1091 | 0.02 | 0.26 |  | 0.32808
589.889 | 0.16014 | 0.041 | 0.054653 | 0.086737 | 0.02 | 0.26 |  | 0.32533
594.465 | 0.16029 | 0.041 | 0.025027 | 0.11359 | 0.02 | 0.26 |  | 0.33
599.448 | 0.16036 | 0.041 | 0.0030562 | 0.11559 | 0.02 | 0.26 |  | 0.3298
604.696 | 0.1588 | 0.041 | 0.0074314 | 0.084417 | 0.02 | 0.26 |  | 0.3195
608.963 | 0.15876 | 0.041 | 0.013159 | 0.07517 | 0.02 | 0.26 |  | 0.31735
615.635 | 0.1588 | 0.041 | 0.013037 | 0.10304 | 0.02 | 0.26 |  | 0.32509
621.063 | 0.15881 | 0.041 | 0.019804 | 0.08159 | 0.02 | 0.26 |  | 0.3193
625.067 | 0.15892 | 0.041 | 0.014715 | 0.065488 | 0.02 | 0.26 |  | 0.31534
631.416 | 0.15885 | 0.041 | 0.0085648 | 0.075779 | 0.02 | 0.26 |  | 0.31738
636.66 | 0.1592 | 0.041 | 0.014131 | 0.069933 | 0.02 | 0.26 |  | 0.31641
644.216 | 0.15883 | 0.041 | 0.01902 | 0.089044 | 0.02 | 0.26 |  | 0.32125
650.776 | 0.15897 | 0.041 | 0.010634 | 0.066072 | 0.02 | 0.26 |  | 0.31533
656.026 | 0.15919 | 0.041 | 0.012112 | 0.083232 | 0.02 | 0.26 |  | 0.31953
659.901 | 0.15902 | 0.041 | 0.010864 | 0.090591 | 0.02 | 0.26 |  | 0.32139
665.942 | 0.1589 | 0.041 | 0.022132 | 0.063693 | 0.02 | 0.26 |  | 0.3154
669.428 | 0.15901 | 0.041 | 0.033881 | 0.061271 | 0.02 | 0.26 |  | 0.31602
674.758 | 0.15935 | 0.041 | 0.0085512 | 0.080449 | 0.02 | 0.26 |  | 0.31878
679.318 | 0.15926 | 0.041 | 0.016341 | 0.055982 | 0.02 | 0.26 |  | 0.31376
684.039 | 0.15969 | 0.041 | 0.049197 | 0.089963 | 0.02 | 0.26 |  | 0.32511
688.896 | 0.15912 | 0.041 | 0.065749 | 0.07038 | 0.02 | 0.26 |  | 0.32292
695.123 | 0.15975 | 0.041 | 0.059707 | 0.1247 | 0.02 | 0.26 |  | 0.33811
699.478 | 0.16063 | 0.041 | 0.095054 | 0.10755 | 0.02 | 0.26 |  | 0.34071
703.605 | 0.15924 | 0.041 | 0.026564 | 0.097272 | 0.02 | 0.26 |  | 0.32435
709.729 | 0.16022 | 0.041 | 0.017556 | 0.1038 | 0.02 | 0.26 |  | 0.32625
715.193 | 0.15988 | 0.041 | 0.022817 | 0.11281 | 0.02 | 0.26 |  | 0.32937
718.438 | 0.16015 | 0.041 | 0.014897 | 0.10939 | 0.02 | 0.26 |  | 0.3279
724.801 | 0.161 | 0.041 | 0.018853 | 0.11194 | 0.02 | 0.26 |  | 0.32938
730.21 | 0.16036 | 0.041 | 0.0094494 | 0.10804 | 0.02 | 0.26 |  | 0.32735
730.5911 | 0.15284 | 0.041 | 0.0094494 | 0.10804 | 0.02 | 0.26 |  | 0.32373
734.9336 | 0.15308 | 0.041 | 0.030779 | 0.11155 | 0.02 | 0.26 |  | 0.32635
737.443 | 0.15926 | 0.041 | 0.019264 | 0.11207 | 0.02 | 0.26 |  | 0.3286
740.0178 | 0.15271 | 0.041 | 0.0071446 | 0.1131 | 0.02 | 0.26 |  | 0.32533
745.512 | 0.15263 | 0.041 | 0.0078674 | 0.12952 | 0.02 | 0.26 |  | 0.33138
750.3849 | 0.15299 | 0.041 | 0.014406 | 0.11865 | 0.02 | 0.26 |  | 0.32768
755.6834 | 0.15269 | 0.041 | 0.013465 | 0.10331 | 0.02 | 0.26 |  | 0.32225
760.3326 | 0.15274 | 0.041 | 0.016562 | 0.11132 | 0.02 | 0.26 |  | 0.32508
764.6821 | 0.15262 | 0.041 | 0.0073396 | 0.10836 | 0.02 | 0.26 |  | 0.32368
770.1808 | 0.15268 | 0.041 | 0.013542 | 0.10195 | 0.02 | 0.26 |  | 0.32182
774.6374 | 0.15268 | 0.041 | 0.013327 | 0.12088 | 0.02 | 0.26 |  | 0.3283
780.2826 | 0.15269 | 0.041 | 0.011301 | 0.12293 | 0.02 | 0.26 |  | 0.32899
785.6348 | 0.15265 | 0.041 | 0.0083528 | 0.11252 | 0.02 | 0.26 |  | 0.32514
790.2739 | 0.15278 | 0.041 | 0.0048086 | 0.12948 | 0.02 | 0.26 |  | 0.33138
795.3709 | 0.15259 | 0.041 | 0.011468 | 0.12902 | 0.02 | 0.26 |  | 0.33127
800.1307 | 0.15279 | 0.041 | 0.010949 | 0.10414 | 0.02 | 0.26 |  | 0.32248
804.9509 | 0.15266 | 0.041 | 0.0048056 | 0.10167 | 0.02 | 0.26 |  | 0.32147
809.8248 | 0.15274 | 0.041 | 0.008155 | 0.11964 | 0.02 | 0.26 |  | 0.3277
815.184 | 0.1527 | 0.041 | 0.0087414 | 0.12318 | 0.02 | 0.26 |  | 0.32901
820.1783 | 0.15346 | 0.041 | 0.0095434 | 0.11624 | 0.02 | 0.26 |  | 0.32685
824.8237 | 0.15288 | 0.041 | 0.0043714 | 0.11977 | 0.02 | 0.26 |  | 0.32774
830.3756 | 0.15303 | 0.041 | 0.013692 | 0.12086 | 0.02 | 0.26 |  | 0.32847
834.6986 | 0.15263 | 0.041 | 0.0091934 | 0.13092 | 0.02 | 0.26 |  | 0.33196
839.9517 | 0.15262 | 0.041 | 0.0067734 | 0.11503 | 0.02 | 0.26 |  | 0.32596
844.3801 | 0.15261 | 0.041 | 0.0039882 | 0.11714 | 0.02 | 0.26 |  | 0.32667
850.6482 | 0.15274 | 0.041 | 0.0019992 | 0.10928 | 0.02 | 0.26 |  | 0.32397
855.6503 | 0.15272 | 0.041 | 0.020162 | 0.11996 | 0.02 | 0.26 |  | 0.32833
860.2478 | 0.15292 | 0.041 | 0.012444 | 0.10605 | 0.02 | 0.26 |  | 0.32321
864.8997 | 0.1526 | 0.041 | 0.005999 | 0.11371 | 0.02 | 0.26 |  | 0.32547
870.06 | 0.15259 | 0.041 | 0.007473 | 0.13204 | 0.02 | 0.26 |  | 0.33235
875.2847 | 0.15263 | 0.041 | 0.010171 | 0.10599 | 0.02 | 0.26 |  | 0.32298
880.5797 | 0.15271 | 0.041 | 0.0059054 | 0.10766 | 0.02 | 0.26 |  | 0.32346
884.4715 | 0.15271 | 0.041 | 0.0054826 | 0.11689 | 0.02 | 0.26 |  | 0.32664
890.3602 | 0.15267 | 0.041 | 0.042817 | 0.10999 | 0.02 | 0.26 |  | 0.32698
895.3368 | 0.15263 | 0.041 | 0.016065 | 0.10862 | 0.02 | 0.26 |  | 0.32409
899.3576 | 0.15264 | 0.041 | 0.046079 | 0.12499 | 0.02 | 0.26 |  | 0.33276
905.4571 | 0.1536 | 0.041 | 0.010248 | 0.09822 | 0.02 | 0.26 |  | 0.32098
910.6081 | 0.15375 | 0.041 | 0.015043 | 0.10141 | 0.02 | 0.26 |  | 0.32222
914.7723 | 0.1555 | 0.041 | 0.015682 | 0.11528 | 0.02 | 0.26 |  | 0.32772
920.56 | 0.16174 | 0.041 | 0.012755 | 0.099628 | 0.02 | 0.26 |  | 0.32547
924.29 | 0.1651 | 0.041 | 0.010864 | 0.10178 | 0.02 | 0.26 |  | 0.32774
930.1835 | 0.17107 | 0.041 | 0.013479 | 0.096829 | 0.02 | 0.26 |  | 0.3294
935.1037 | 0.16394 | 0.041 | 0.032675 | 0.094408 | 0.02 | 0.26 |  | 0.3264
940.6012 | 0.16618 | 0.041 | 0.019472 | 0.10892 | 0.02 | 0.26 |  | 0.33097
944.4859 | 0.17092 | 0.041 | 0.017421 | 0.090538 | 0.02 | 0.26 |  | 0.32771
950.6383 | 0.16972 | 0.041 | 0.002312 | 0.053603 | 0.02 | 0.26 |  | 0.31838